## Slide 1: Evaluation of Intelligent Phaco Mode in Torsional Phacoemulsification
David Allen
Sunderland UK
Financial disclosure A & E
  The author has had travel and lodging costs paid by as well as occasional honoraria from Alcon Surgical

## Slide 2: Background
Torsional phaco works by shaving material. If a lump of nucleus becomes lodged in the opening of a non-optimum flared or mini-flared tip it can stall the consumption or even cause clogging
I have shown (ESCRS annual meeting, Stockholm 2007) that this problem is minimised by using the 45° mini-flare Kelman tip
However, not all surgeons are comfortable with this tip

## Slide 3: Background
IP is a new mode which has been introduced on the Infiniti OZil platform, designed to reduce or eliminate this potential problem when surgeons use other tips/needles
When the vacuum level approaches or reaches the preset maximum, short (10ms) pulses of longitudinal phaco are added every 100 ms until the vacuum drops
The theory is that the traditional longitudinal pulses keep material at the shearing plane ensuring emulsification before it enters narrowing part of the flare or mini-flare tip

## Slide 4: Questions and methods
If additional longitudinal phaco is being used, is there an ‘energy penalty’?
Does it add to the smoothness of the surgery?
Series of studies (50 cataracts in each series) using ‘stop and chop’ to divide the nucleus in two halves
Each half is removed with the different settings being compared (therefore each acts as its own control for nucleus density)

## Slide 5: Results – CDE measured with IP on and IP off
Mini-flare Kelman needle used throughout
‘moderate fluidics’ = 300 mmHg vacuum limit and 30 mL/min aspiration flow rate
‘high fluidics’ = 400 mmHg vacuum limit and 40 mL/min aspiration flow rate
No significant difference in CDE between IP on/off in any of the groups (i.e. NO power penalty)
 | IP On | IP Off
45° tip and moderate fluidics | 2.70 | 2.69
30° tip and moderate fluidics | 2.56 | 2.85
30° tip and high fluidics | 2.11 | 2.24

## Slide 6: Results – procedure ‘smoothness’ with IP on and IP off with 30° mini-flare Kelman tip
 | Stalls - >12 video frames | Clogs
30° moderate fluidics IP on | 3 | 0
30° moderate fluidics IP off | 26 | 0
30° higher fluidics IP on | 4 | 0
30° high fluidics IP off | 31 | 4
Analysis of recorded video to find number of occasions when the procedure ‘stalls’ – i.e. maximum vacuum reached but no apparent material consumption for >12 video frames (500 ms)
‘Clog’ defined as full occlusion for >2 seconds despite maximum ‘power’

## Slide 7: Part 2 of study
These results confirmed the clinical impression that IP is effective and that most benefit is seen with 30° mini-flare needle.
The next stage was to see what effect different settings might have with that needle
Assessed the CDE value for each hemi-nucleus, and the number of ‘stalls’ of the procedure - this time using a more sensitive threshold of > 5 consecutive video frames (200 ms)

## Slide 8: Part 2 of study
Examine the threshold for IP being activated – either at 95% of target vacuum or only when full vacuum reached (100%)
Examine how much longitudinal phaco added – ratio of longitudinal power applied to how much torsional amplitude being commanded by the footpedal

## Slide 9: Different settings of OZil IP – vacuum threshold
 | CDE | Stalls - >5 video frames
Trigger vacuum threshold 95% | 2.61* | 24 *
Trigger vacuum threshold 100% | 2.38* P=0.12 | 37 * P=0.02
There is a small but statistically insignificant reduction in CDE if the IP does not trigger until the max. vacuum is reached
However there is a statistically very significant decrease in the number of ‘stalls’ seen on the video with vacuum threshold set at 95%

## Slide 10: Different settings of OZil IP – vacuum threshold
 | CDE | Stalls - >5 video frames
Long. power to OZil amplitude ratio 0.7 | 2.08 | 36
Long. power to OZil amplitude ratio 1.0 | 2.07 | 31
The ratios chosen were the maximum and minimum allowed by the software
The ratio of longitudinal power to torsional amplitude commanded (within the available range) does not appear to have an impact on the CDE used or the smoothness of the procedure

## Slide 11: OZil IP summary (1)
IP appears to have no energy use downside
IP clearly makes the procedure smoother with fewer interruptions when using less than optimal tips.
The greatest benefit is seen with a vacuum threshold set at 95% of the maximum
The ratio of the added longitudinal power to commanded torsional amplitude seems to have little or no impact on the smoothness or total phaco power.

## Slide 12: OZil IP summary (2)
This study was performed with moderate to hard cataracts as seen in Sunderland (UK)
It is reasonable to suppose that IP would have even more benefit in regions where ‘rock’ cataracts are more common